Project Name: Humidity Data Logger
Company: Innovation Device Solution (IDS)
Version: V0.0
Date: 23-09-2025

1. General Information

| Parameter | Requirement |
| --- | --- |
| Device Type | Humidity & Temperature Data Logger |
| End Applications | Environmental monitoring, Industrial process, Agriculture, Storage/Cold Chain |
| Enclosure Rating | IP54 minimum |
| Compliance | CE / RoHS / EMC (IEC 61000-4-2/3) |

2. Sensor Specifications

| Parameter | Requirement |
| --- | --- |
| Sensor Model | TI HDC3022 |
| Humidity Range | 0–100 %RH |
| Humidity Accuracy | ±2 %RH (20–80 %RH) |
| Humidity Resolution | 0.01 %RH |
| Temperature Range | -20 °C to +85 °C |
| Temperature Accuracy | ±0.5 °C (0–60 °C) |
| Temperature Resolution | 0.01 °C |
| Step Response (τ) | 4 s (63% RH step, moving air) |
| Long-Term Drift | ≤0.19 %RH/year |

3. Logging & Storage

| Parameter | Requirement |
| --- | --- |
| Logging Interval | 5s – 1 hr (configurable) |
| Memory Capacity | ≥64,000 readings |
| Data Type | Humidity, Temperature, Timestamp |
| RTC | Integrated with battery backup |

4. Power & Battery

| Parameter | Requirement |
| --- | --- |
| Battery Type | CR2450/2032 Coin Cell |
| Typical Battery Life | ≥1 year @ 15 min logging interval |
| Indicators | Low Battery Warning |

5. User Interface

| Parameter | Requirement |
| --- | --- |
| Display | LCD 3 ½ showing RH, Temp, Status |
| Buttons | 2 Keys (Start/Stop, Min/Max) |
| LED Indicators | Logging, Battery Low |

6. Connectivity

| Parameter | Requirement |
| --- | --- |
| USB | USB 2.0 (data transfer & config) |
| Data Export | CSV / PDF via PC Software |

7. Software

| Parameter | Requirement |
| --- | --- |
| IDS Logger PC Software | Configuration, Data Download, Visualization |

8. Environmental Conditions

| Parameter | Requirement |
| --- | --- |
| Operating Temp | -20 °C to +60 °C |
| Operating Humidity | 0–95 %RH (non-condensing) |
| Storage Conditions | -30 °C to +70 °C |